| Designation | Desinfectant Wipes Medibase |
| --- | --- |
| Reference(s) | The box of 100 wipes | 3131820 |
|  | The refill bag of 100 wipes | 3131821 |
| Product picture |  |
| Product description | Wipes impregnated with a disinfectant solution for cleaning and disinfecting non-invasive and non-submersible medical devices (small medical and dental equipment). |
| Characteristics and compositon | Less than 5% non-ionic surfactants, perfumes including limonene, disinfectants: Ethanol, Alkyl, Dimethylbenzylammonium chloride, Propan-2-ol. Wipes impregnated with a disinfectant solution for cleaning and disinfecting non-invasive and non-submersible medical devices (small medical and dental equipment). - Wipes particularly large, thick, strong and sweet. - Less than 5% non-ionic surfactants, perfumes including limonene, disinfectants: Ethanol, Alkyl, Dimethylbenzylammonium chloride, Propan-2-ol. ACTIVITE MICROBIENNE Contact time General treatments Bactericidal EN13727* - EN14561 5min EN14561: MRSA 5min Fungicidal EN14562: Aspergillus fumigatus 5min Yeasticidal EN13624 - EN 14562: Candida albicans 5min Tuberculocidal EN14348 - EN 14563 : Mycobacterium terrae 15min Virucidal EN14476* : HBV - HCV - Adenovirus Coronavirus - Novovirus - VRS - H1N1 - HSV1 1min EN14476: Polyomavirus 1min EN14476: Rotavirus 15min Additional uses Bactericidal EN1276 30s EN13697 5min Fungicidal EN13697: Aspergillus niger 5min Yeasticidal EN1650* - EN 13697: Candida albicans 5min *In dirty conditions |
| Packaging | - The box of 100 wipes The refill bag of 100 wipes |
| Shelf life | 3 month from the day of opening |
| Indications of use | Ready-to-use wipes. Single-use wipes. When using the box for the first time: Remove the cover, Permanently remove the seal, Thread the central wipe through the hole in the cover with clean hands and Replace the cover. Use the wipe on the equipment. Leave on for 30 seconds to 15 minutes according to the antibacterial activity required. Leave to dry in the open air. Do not rinse. Close the flap after each use to prevent the wipes from drying out. |
| Legal information | Medical Device of class IIA CE Marking 0459. G-MED Notified body. Manufacturer : FELT S.A |

Version 2.2						Updated 15/11/2022 by RG

# Text boxes

TECHNICAL SHEET